*** Data and Peak File Formats

There are two types of data files recognized in the Powder Suite

programs, data files and peak files. Data files usually contain the

raw diffraction results and values describing how the measurement was made,

while peak files contain derived results such as peak positions, heights and

areas. Occasionally a single file may contain both sets of information.

The philosophy used in the Powder Suite programs is to allow the programs

to read as many different types of data files as can be conveniently

incorporated and to allow the programs to be easily expanded to accept

additional formats.

Also, rather than require users to specify the format of the file

they want to read, the code has been written to automatically determine

the format by reading the file.

Associated with each file format is a default extension.

For example, DuPont files are expected to have extension .DDAT,

and Brookhaven X7A files are expected to have extension .X7A.

If a program is instructed to read data file TEST,

then it will first search for a file named TEST.DDAT and if it is found

check if it is in DuPont format, if no such file is found or the file cannot

be read, it will then search for a file named TEST.X7A.

Successive attempts will be made until each file type has been tried.

If the file name is supplied to the program with an extension, for example

TEST.XXX or TEST.DDAT, the procedure used is

slightly different. In this case, the program will attempt to read the

file successively in each format until it is successfully read.

On occasion, it may be necessary to specify the file format to be used.

This may be done by following the file name with a /FORMAT= and

the file format: filename[.ext]/FORMAT=DUPONT. When the filename is

specified on the command line, the /FORMAT= qualifier may be abbreviated,

but must be used. However, if the filename is specified in response to

a prompt from the program, the string FORMAT= may be omitted:

filename.ext/DUPONT.

Dupont.

Dupont binary format files contain diffractogram data and

optionally peaks and baseline points.

The default file extension is .DDAT; FORMAT=DUPONT.

NIST (NBS).

The NIST format is a simple ASCII format for raw

intensity data. The file contains two header lines

followed by intensity values as either counts per

second or counts, ten per line. The count time is not

stored, so ESD's are computed assuming the intensity values

are stored as counts.

The default file extension is .NBS; FORMAT=NIST.

GSAS.

The programs can read some of the ASCII GSAS Standard

Powder Data files used as input to GSAS (BINTYP='CONST',

optionally with ESD's). See the GSAS manual for details.

The default file extension is .GSAS; FORMAT=GSASASC.

In addition, information can also be read from GSAS

histogram and reflection files (refinement output).

The program must read several files, including the .EXP,

the .REFxxP file and the .Pxx file.

The default file extension is .EXP; FORMAT=GSASHIST.

BNLX7A.

This is the format used for data recorded at the

Brookhaven NSLS X7A (formerly X13A) instrument.

The source flux changes with ring current so

intensities are recorded with both scaler and monitor

values versus either Q or two-theta. After reading,

Intensities are scaled to constant monitor counts.

The default file extension is .X7A; FORMAT=BNLX7A.

JONES.

This highly complex format was developed by Glover

Jones of DuPont. It consists of one angle and one

intensity value (format 2F10.3) per line, with no

header information. Data need not be recorded with equal

steps of two-theta, but the first two lines must have

decimal points (.) in columns 7 and 17.

The default file extension is .JONES; FORMAT=JONES.

IPNS.

Binary files produced by the Penn versions of Argonne

Intense Pulsed Neutron Source (IPNS) GLASS package and

TOF package may be read: the file types that can be read are

Radial Distribution Functions, rho(r), g(r), n(r), (default file

extension .RDF, FORMAT=IPNSRHO, IPNSGOFR, IPNSNOFR);

Structure factors, S(Q), S(d) (default file extension .SOQ,

FORMAT=IPNSSOFQ, IPNSSOFD)

ASCII raw counts from .RUN files written by RUNPLT

(default file extension .TOF, FORMAT=IPNSTOFx where

x = T for time, L for wavelength, Q for Q and D for d-spaces)

Argonne TOF Rietveld .PLT files written by TOFLS

(default file extension .PLT, FORMAT=IPNSRIETxy where

x = T for time, L for wavelength, Q for Q and D for d-spaces

and y = O for the observed, background subracted pattern,

C for the computed pattern and D differences [obs-calc]).

PHILIPS -- Nova.

This format corresponds to an ASCII data file created by

the Philips APD system running on a Data General Nova

computer. The file contains a single header line with the

step size, the starting and the ending angle, followed by

intensity values as either counts per second or counts,

ten per line. The count time is not stored, so ESD's are

computed assuming values in counts. The source radiation

is assumed to be both Cu K alpha 1.

The default file extension is .RD; FORMAT=APDNOVA.

PHILIPS -- VAX.

This format corresponds to an binary data file created by

the Philips 3720 APD system running on a VAX/VMS

computer and two associated peak files, one ASCII and

one binary. The binary data file consists of fixed length,

512 byte records and the default file extension is .RD.

The binary format peak file also has 512 byte records and

has a default extension of .DI. The ASCII peak file is

the contents of a peak report from the APD 3720 and

has a default extension of .APLIS;

The FORMAT is APDVAX for all of these files.

SCINTAG -- VAX.

This format corresponds to an binary data file created by

the Scintag system running on a VAX/VMS and two associated

peak files. The data file consists of fixed length, 4 byte

records and the default file extension is .RD. The binary

format peak file also has 4 byte records and has default file

extension is .PK. The ASCII format peak file has default

extension .PKLST. Since the x-ray tube type is not stored,

the radiation is assumed to be both Cu K alpha 1

and K alpha 2. FORMAT=SCINTAG.

Penn State University "ASCII data file" and "Standard D-I format"

peak files.

The data file has a title for the first line. The second line

contains the starting, ending angles, step size and count time

(Format 4F10.0). The remainder of the file has 10 intensity

values (as counts) per line (Format 10F8.0).

The default data file extension is .ASC.

The peak file has any number of header lines, followed by a

card with "PEAK" in the first four columns.

Remaining lines have one d-space and one peak height per line

(Format 14X,F8.4,2X,F3.0).

The default peak file extension is .ADR.

In addition to the peak files associated with data files, several

additional peak files may be read. They all use FORMAT=PEAK.

SEARCH temporary (peak) files.

The temporary files with lists of peak positions may be

read by other programs as a peak file. For example, if peaks

are entered into SEARCH for use in identification of an

unknown pattern, the same peak list may be displayed along

with peak positions from JCPDS-ICDD cards in program LOGIC.

The default file extension is .SEA_IN.

Free Format peak file.

The first line in this file is read as a title and the second

line (format F10.0) is a wavelength. The wavelength should be

zero if peak positions have been entered as d-spaces and non-zero

for two-theta angles. Remaining lines have one

d-space and one peak height per line (Format 2F10.0). Note

that positions and heights may be separated by a comma, if

the position and both preceeding and following spaces are less

than 10 characters. For example

10.01, 99.

is valid.

These files must have extension .PEAK to be read.